## Schematic
Component | Description | Proposed control
V1 | Steam master on/off. | Technician controlled
R1 | Pressure regulator used to control steam volume to working fluid. | Technician controlled
M | Steam pressure gauge | Add pressure sensor - Gems?
V2 | Not used-always fully on. | NA
V3 | Used to throttle the steam to working fluid-adjusted by watching condensation on secondary condesnding coils (2-3 coils). | replace with Burkert motorised needle valve if necessary, ideally "set and forget"
V4 | Needle valve integrated on flow meter - used to control cooling on primary condensor. | leave fully open, add Enfield valve inline to control water flow, add RS digital flowmeter
V5 | Fully open - cooling unit turned on or off only. | Technician controlled, perhaps remote controll on chiller unit
V6 | Used to drain working fluid in conjunction with V8. | Technician controlled
V7 | Used to dry rig between working fluid changes in conjunction with V8. | Technician controlled
V8 | Used with V6 and V7 above. | Technician controlled
TIN | Primary condensor inlet water temp. | Tap into thermocouple signal
TOUT | Primary condensor outlet water temp. | Tap into thermocouple signal
TV | Primary condensor environment temperature | Tap into thermocouple signal
TL | Working fluid temperature | Tap into thermocouple signal

## System IOs
| Part required | Manufacturer | Mfr P.N. | Price (£) | Price € | Qty Rq'd | Total | Control/Indicator | Control/Signal Type | Signal Range | Notes |  | Additional info | Supplier | Delivery time | Link
| Water flow meter | RS | 257-149 | 55.92 |  | 1 | 55.92 | indicator | Pulse | 4.5-16 VDC | Assume Vout = Vin |  | For digital output of cold water flow | RS | in stock | https://uk.rs-online.com/web/p/flow-sensors/0257149
| Motorised needle valve | Enfield | ENV-0670 | 470 |  | 1 | 470 | control | Driver Board (see below) |  |  |  | For controlling cold water flow 577usd | Enfield | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/ENV-0670-Motorized-Needle-Valve
| Driver for valve | Enfield | D5-05-U01 | 180 |  | 1 | 180 | control | Analog Voltage | 0-10v |  |  | As above 220USD | Enfield | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/D5-05-U01
| Pressure sensor | Gems |  |  |  | 1 |  | control | Analog Voltage? | 0-10v? |  |  | emailed Gems for advice 31/5/23 | Gems sensors |  | https://www.gemssensors.co.uk/
| Proportional steam valve | Burkert | 375565 | 3716 |  | 1 | 3716 | [5 empty] | provisional quote based on rough estimate of required flow rate | Burkert | 9-10 weeks | https://www.burkert.co.uk/en
[6 empty] | Total | 4421.92
[6 empty] | 20% | 5306.304 | [5 empty] | Additional info | Supplier | Delivery time | Link
| Parts required - ELECTRIC HEATING | Manufacturer | Mfr P.N. | Price |  | Qty Rq'd | Total | [5 empty] | For digital output of cold water flow | RS | in stock | https://uk.rs-online.com/web/p/flow-sensors/0257149
| Cold Water flow meter | RS | 257-149 | 55.92 |  | 1 | 55.92 | indicator | Pulse | 4.5-16 VDC | Assume Vout = Vin |  | For controlling cold water flow 577usd | Enfield | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/ENV-0670-Motorized-Needle-Valve
| Motorised needle valve | Enfield | ENV-0670 | 470 |  | 1 | 470 | control | Driver Board (see below) |  |  |  | As above 220USD | Enfield | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/D5-05-U01
| Driver for valve | Enfield | D5-05-U01 | 180 |  | 1 | 180 | control | Analog Voltage | 0-10v
| Heating Element |  |  | 300 |  | 1 | 300 | control | Relay?
[6 empty] | Total | 1005.92
[6 empty] | 20% | 1207.104
Requirement | Electronic Equipment Specified | Manufacturer | Mfr P.N. | Price |  | Qty Rq'd | Total | Control/Indicator | Control/Signal Type | Signal Range |  |  | Additional info | Supplier | Delivery time | Link
Cold Water flow meter | ADAM-5080 | Advantech | ADAM-5080 | 320 | ESTIMATED | 1 | 320 | Indicator | Pulse | 5v TTL |  |  |  | Advantech |  | https://www.advantech.com/th-th/products/a85377de-9600-4649-a0d6-b0c84ef3202d/adam-5080/mod_95ed1d56-38b2-4ac7-86d1-63abfed56e2c?gclid=Cj0KCQjwho-lBhC_ARIsAMpgMofQAaDFswcUe4qLYRmo2Hm-DHaCOrz1zZa9FAKrvvdOwnGBxi7reHUaAg9gEALw_wcB
| 4-ch Counter/Frequency Module |  | "get quote"
Needle Valve Driver | ADAM-5024-A2E | Advantech | ADAM-5024-A2E |  | 327.7 | 1 | 327.7 | Control | Analog Voltage/Current | 0-10V (+ 0-20mA) |  |  |  | Advantech Online | "Available for Delivery" | https://buy.advantech.eu/I-O-Devices-Communication/Remote-I-O-Modules-ADAM-I-O-Modules/model-ADAM-5024-A2E.htm
Heating Element Relay | ADAM-5056D-AE | Advantech | ADAM-5056D-AE |  | 105.2 | 1 | 105.2 | Control | Digital | Open Collector to 30v |  |  |  | Advantech Online | "contact for stock" | https://buy.advantech.eu/I-O-Devices-Communication/Remote-I-O-Modules-ADAM-I-O-Modules/model-ADAM-5056D-AE.htm
| AND | [8 empty] | 100mA max
Heating Element Relay | Finder DIN Rail Mount Interface Relay, 12V dc Coil, 16A Load Current, SPDT | Finder | 48.P6.7.012.0050SPA | 13.14 |  | 1 | 13.14 | Control | Relay | 240v switching 16A load |  |  |  | RS online | in stock | https://uk.rs-online.com/web/p/electromechanical-interface-relays/1232816?gb=s
12v coil
DAQ Chassis | ADAM 4 Slot Ethernet |  |  |  | 398 | 1 | 398 |  |  | 10-30v |  |  |  | Advantech |  | https://buy.advantech.eu/I-O-Devices-Communication/Remote-I-O-Modules-ADAM-Controllers/model-ADAM-5000L/TCP-BE.htm
| ADAM-5000 DAQ Chassis | [12 empty] | Impulse-Embedded |  | https://www.impulse-embedded.co.uk/products/adam_5000l~tcp--Distributed-DAQ-system.htm
PSU | RS PRO DIN Rail Power Supply, 85 → 264 V ac, 120 → 370 V dc ac, dc Input, 12V dc Output, 10A Output, 120W |  |  | 86.28 |  | 1 | 86.28 | [6 empty] | RS | in stock | https://uk.rs-online.com/web/p/din-rail-power-supplies/2205404?gb=s
Enclosure | [5 empty] | 1 | 150
Sunds Electronics (terminals etc) | [5 empty] | 1 | 150
[10 empty] | Electric |  | Steam
[5 empty] | Electronics |  |  |  | Total Total | 3067.488 |  | 7166.688
[5 empty] | Total |  | 1550.32
[5 empty] | 20% |  | 1860.384

## Mechanical parts
Part required | Manufacturer | Mfr P.N. | Price | Supplier | Qty Rq'd | Total | Delivery time | Link | Additional info
Water flow meter | RS | 257-149 | 55.92 | RS | 1 | 55.92 | in stock | https://uk.rs-online.com/web/p/flow-sensors/0257149 | For digital output of cold water flow
Motorised needle valve | Enfield | ENV-0670 | 470 | Enfield | 1 | 470 | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/ENV-0670-Motorized-Needle-Valve | For controlling cold water flow 577usd
Driver for valve | Enfield | D5-05-U01 | 180 | Enfield | 1 | 180 | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/D5-05-U01 | As above 220USD
Pressure sensor | Gems |  |  | Gems sensors | 1 |  |  | https://www.gemssensors.co.uk/ | emailed Gems for advice 31/5/23
Proportional steam valve | Burkert | 375565 | 3716 | Burkert | 1 | 3716 | 9-10 weeks | https://www.burkert.co.uk/en | provisional quote based on rough estimate of required flow rate
Parts required - ELECTRIC HEATING | Manufacturer | Mfr P.N. | Price | Supplier | Qty Rq'd | Total | Delivery time | Link | Additional info
Cold Water flow meter | RS | 257-149 | 55.92 | RS | 1 | 55.92 | in stock | https://uk.rs-online.com/web/p/flow-sensors/0257149 | For digital output of cold water flow
Motorised needle valve | Enfield | ENV-0670 | 470 | Enfield | 1 | 470 | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/ENV-0670-Motorized-Needle-Valve | For controlling cold water flow 577usd
Driver for valve | Enfield | D5-05-U01 | 180 | Enfield | 1 | 180 | in stock | https://www.enfieldtech.com/Products/ENV-2-2-MOTORIZED-NEEDLE-VALVES/D5-05-U01 | As above 220USD